VALIDATION REPORT CAELYX®

HIGH LEVEL

PROTOCOL No. UMCNONCO200902

- SUMMARY VALIDATION PERFORMANCES

The following performances  are tested: Selectivity (co-medication and interference endogenous components matrix), Stability autoinjector, Caelyx depegylation within-run (dispersion test), Caelyx efficiency depegylation between-run (dispersion test), Between-run doxorubicine/doxorubicinol, Within-run doxorubicine/doxorubicinol, Freeze-thaw stability, 24 hour room temperature stability, Recovery, Stability longterm ( pending )

- ASSAY PROCEDURE

Analytes: 		Doxorubicin, Doxorubicinol.

Internal standard: 	Daunorubicin

  - 250 μL Standard/Blank/QC/Sample + 25 μL Triton 6% (v/v) in eppendorf®-cup.

  - Mix 2 minutes on vortex

  - Wait 10 minutes

  - Add 250 μL internal standard mix (IS + deproteinization solution). Except for blank( 250 μL deproteinization solution without IS

  - Mix 2 minutes on vortex

  - Centrifuge 5 minutes at 10900 RPM

  - Transfer 250 μL in glas vial

  - Inject 5 μL onto HPLC-system

- INSTRUMENTATION AND CONDITIONS

HPLC-system: 	Shimadzu LC-20AT HPLC-pump with DGU 20A3-degasser

Waters 717plus autosampler

Jasco FP-920 Intelligent Fluorescence detector

HPLC-column:	Waters Atlantis T3, 15 x 4.6 mm; 5 μm

Guard-column:	Phenomenex Security Guard C18 4 x 3.0 mm

Mobile phase:		Acetonitril : Sodiumacetate buffer 25:75

Needlewash:		Acetonitril : Water 25:75

Flow:			2,0 ml/min

Excitation λ:		480 nm

Emmission λ:		560 nm

Bandwith:		40 nm

Gain:			100

Temperature:		Ambient

Runtime:		22 minutes

Injection volume:	5 μL

VALIDATION REPORT CAELYX®

HIGH LEVEL

PROTOCOL No. UMCNONCO200902

- APPLIED ANALYTICAL METHOD

Own method, based on articles Peter de Bruijn et al: Determination of doxorubicin en doxorubicinol in plasma of cancer patients by high-performance liquid chromatography ( Analytical Biochemistry 266, 216-221, 1999 ) AND David L. Chin et al: Rapid determination of PEGylated liposomal doxorubicin and its major metabolite in human plasma by ultraviolet-visible high-performance liquid chromatography ( Journal of Chromatography B, 779 259-269, 2002 )

- REAGENTS / MOBILE PHASE

Sodiumacetate-buffer:6,8 grams / liter water ( Na-Ac. 3 H2O ). pH=4.0 with Acetic acid

Triton 6% (v/v): 6 ml Triton X-100 in 100 ml water. Heat to 40 °C to dissolve.

Zinksulfate 89g/L: 89 grams of  ZnS04.7H20 in liter water

Deproteinization solution: Zinksulfate 89g/L : Methanol ( 1: 1 )

Internal standard MIX: add 150 μL for every 3 mL of deproteinization solution

- REFERENCE STANDARDS

| COMPONENT | ORIGIN | BATCH NR | CERTIFICATE | STABILITY | STORAGE CONDITIONS |
| --- | --- | --- | --- | --- | --- |
| Doxorubicin | LC labs | DXR-102 | YES | 18-1-2016 | -20 °C |
| Doxorubicinol | Alsachim | CM-ALS-11-088-P3 | YES | 21-9-2013 | -20 °C |
| Daunorubicin | Sigma | 1352629V | NO | ? | 5 °C |

- CALIBRATION STANDARDS AND QC SAMPLES

Standard solution

| COMPONENT | WEIGHT FACTOR | C | MATRIX | PREPARATION | PREPERATION DATE | STORED |
| --- | --- | --- | --- | --- | --- | --- |
| Daunorubicin | .1 HCl | 1000 mg/L | Methanol | 9.77 mg in 10 mL | 28-09-2012 | 5 °C |
| Doxorubicin | .1 HCl | 1000 mg/L | Methanol | 10.84 mg in 10 mL | 28-09-2012 | 5 °C |
| Doxorubicinol | .3 Fluoroacetate | 1000 mg/L | Methanol | 12.55 mg in 10 mL | 28-09-2012 | 5 °C |

Standard solution for QC samples

| COMPONENT | WEIGHT FACTOR | C | MATRIX | PREPARATION | PREPERATION DATE | STORED |
| --- | --- | --- | --- | --- | --- | --- |
| Doxorubicin | .1 HCl | 1000 mg/L | Methanol | 10.80 mg in 10 mL | 05-11-2012 | NONE |
| Doxorubicinol | .3 Fluoroacetate | 1000 mg/L | Methanol | 12.32 mg in 10 mL | 05-11-2012 | NONE |

VALIDATION REPORT CAELYX®

HIGH LEVEL

PROTOCOL No. UMCNONCO200902

- RUN ACCEPTANCE CRITERIA

| PARAMETER | CRITERIA | PASSED | REMARK |
| --- | --- | --- | --- |
| Selectivity: Co-medication | No interference | YES | List components(Excel file |
| Selectivity: Blank Matrix | Response <20% of response analytes LLOQ Response <5% of response IS | YES |  |
| LLOQ | Signal >5 times signal blank | YES |  |
| Carry Over | Response blank <20% of response analytes LLOQ Response blank <5% of response IS | YES |  |
| Back Calculated concentrations curve | <15% for QC L,M,H <20 % for QC LLOQ | YES |  |
| Within-run Accuracy | <15% for QC L, M,H <20 % for QC LLOQ | YES |  |
| Between-run Accuracy | <15% for QC L, M,H <20 % for QC LLOQ | YES |  |
| Within-run Precision | VC <15% QC L,M,H VC <20% QC LLOQ | YES |  |
| Between-run Precision | VC <15% QC L,M,H VC <20% QC LLOQ | YES |  |
| Stability autoinjector | <15% of nominal concentration | YES |  |
| Stability Freeze-thaw | <15% of nominal concentration | YES |  |
| Stability long therm | <15% of nominal concentration | PENDING |  |
| Stability 24 hour | <15% of nominal concentration | YES |  |

- ANALYSIS

Analytical runs

| PARAMETER | ANALYSIS DATE | REMARK | PASSED |
| --- | --- | --- | --- |
| Selectivity: Co-medication | 08-11-2012 |  | YES |
| Selectivity: Blank Matrix | 09-11-2012 |  | YES |
| LLOQ | 11/12/13/14-11-2012 |  | YES |
| Carry Over | 14-11-2012 |  | YES |
| Back Calculated concentrations curve | 11/12/13/14-11-2012 |  | YES |
| Within-run Accuracy | 14-11-2012 |  | YES |
| Between-run Accuracy | 11/12/13-11-2012 |  | YES |
| Within-run Precision | 14-11-2012 |  | YES |
| Between-run Precision | 11/12/13-11-2012 |  | YES |
| Stability autoinjector | 11-11-2012 |  | YES |
| Stability Freeze-thaw | 14-11-2012 |  | YES |
| Stability long therm | PENDING |  | YES |
| Stability 24 hour | 14-11-2012 |  | YES |
| Depegylation within-run | 14-11-2012 |  | YES |
| Depegylation between-run | 15/16-11-2012 |  | YES |
| Recovery | 11/12/13-11-2012 | Added(Excel file | NO CRITERIA E.M.A |

VALIDATION REPORT CAELYX®

HIGH LEVEL

PROTOCOL No. UMCNONCO200902

Calibration results

| Calibration date | Range | Response | Back Cal. Concentration | Accuracy |
| --- | --- | --- | --- | --- |
| 15-11-2012 | 0.5 – 50 mg/L DOXO | 5.25 E-02 – 6.87 E-04 | (Excel sheet | (Excel sheet |
| 15-11-2012 | 0.5 – 50 mg/L DOXOL | 1.95 E-02 + 1.77 E-03 | (Excel sheet | (Excel sheet |
| 12-11-2012 | 0.5 – 50 mg/L DOXO | 5.53 E-02 + 6.43 E-05 | (Excel sheet | (Excel sheet |
| 12-11-2012 | 0.5 – 50 mg/L DOXOL | 2.08 E-02 + 1.99 E-03 | (Excel sheet | (Excel sheet |
| 13-11-2012 | 0.5 – 50 mg/L DOXO | 5.43 E-02 + 1.87 E-03 | (Excel sheet | (Excel sheet |
| 13-11-2012 | 0.5 – 50 mg/L DOXOL | 2.01 E-02 + 4.23 E-03 | (Excel sheet | (Excel sheet |
| 14-11-2012 | 0.5 – 50 mg/L DOXO | 5.79 E-02 – 2.23 E-04 | (Excel sheet | (Excel sheet |
| 14-11-2012 | 0.5 – 50 mg/L DOXOL | 2.11 E-02 + 1.65 E-03 | (Excel sheet | (Excel sheet |

Concentrations in plasma ( derived from St 1000 mg/L )

|  | CONC. DOXORUBICIN mg/L | CONC. DOXORUBICINOL mg/L |
| --- | --- | --- |
| Blank | 0 | 0 |
| Standard 0 | 0 | 0 |
| Standard 1 | 0.5 | 0.5 |
| Standard 2 | 1.5 | 1.5 |
| Standard 3 | 5.0 | 5.0 |
| Standard 4 | 10.0 | 10.0 |
| Standard 5 | 25.0 | 25.0 |
| Standard 6 | 50.0 | 50.0 |
| QC 1 ( LLOQ ) | 0.5 | 0.5 |
| QC 2 ( LOW ) | 1.5 | 1.5 |
| QC 3 ( MED ) | 25.0 | 25.0 |
| QC 4 ( HIGH ) | 40.0 | 40.0 |

|  | DOXO WITHIN-RUN | DOXO BETWEEN-RUN | DOXOL WITHIN-RUN | DOXOL BETWEEN-RUN |
| --- | --- | --- | --- | --- |
| QC LLOQ | 95.6% CV 2.9% | 97.8% CV 1.2% | 87.6% CV 3.4% | 92.2% CV 8.7% |
| QC LOW | 92.6% CV 1.5% | 93.3% CV 3.1% | 94.7% CV 2.1% | 98.1% CV 3.5% |
| QC MED | 93.9% CV 2.3% | 96.2% CV 1.3% | 97.3% CV 1.6% | 99.6% CV 0.6% |
| QC HIGH | 96.6% CV 1.0% | 99.3 % CV 2.6% | 98.4% CV 1.7% | 98.8% CV 2.1% |

- SUPPORTIVE DATA

All not summarized data can be found in Excel-file: “Validatie hoge gebied Caelyx studie”

VALIDATION PERFORMED ACCORDING TO: 	European Medicines Agency EMEA/CHMP/EWP/192217/2009

VALIDATION DONE BY: 			Y. de Beer

REPORT DATE: 	 			10-12-2012

AGREED	: 				__________________

DATE: 					__________________